## Slide 1
<number>
DART Tutorial
Data Acquisition in Real Time
This Tutorial is showing a detailed instructions on how to use DART
Notes: <number>
Record A job
DART Tutorial

## Slide 2
<number>
Dart Tutorial
Data Acquisition in Real Time
Upon starting DART the program will start job recording automatically creating file with the following format
EX: 130631.mdb which is the current date in the following format 13/06/31 (Year,Month,day / 2013 June 31)
But if the DART was already opened before on the same day and a file with the same Date has been created, then this window will appear prompting the user whether to “create a new file” for recording or “Append to existing file”
Once the recording start the record button turns from red to green
02:52:02 0.000 0.000 0.000 0.000 1433.000 1433.000 0.000 0.000 0.000 0.0000 24.0000 12.0000
These are the reading coming out of the H.W system it should be available as well as it is connected
Notes: <number>
Record A job
DART Tutorial

## Slide 3
<number>
DART Tutorial
Acquire and record manually
The Following steps shows how To Acquire data and record manually for any reason ex:after stopping recording then recording back
Click Acquire Data from the toolbox
1
Then click on Record button from the toolbox
2
Then this window will appear prompting user to specify a name for mdb file for job recording
3
Note : Make sure that the Record button turns to green, also to make sure that the system is properly connected the Raw value data must be available at the top beside the logo as mentioned in the previous slide
Notes: <number>
Record A job
DART Tutorial

## Slide 4
<number>
DART Tutorial
Acquiring and Recording Data
After specifying an mdb file name this window will appear prompting user to select Parameters to be included in the summary report, these parameters will appear in the summary table.
To include a parameter tick the check box next to the parameter
When you are done click close.
Then recording and plotting of the Data will begin
Notes: <number>
Record A job
DART Tutorial

## Slide 5
<number>
DART Tutorial
Parameters properties
This slide shows how to edit the parameter properties as name, unit, curve color, min. and max. scale
First, right click on the parameter name you wish to edit EX weight indicator
Then, This window will appear, allowing user to input all parameters
Press this button to calibrate
Click here to reset or preset a value to the parameter
Notes: <number>
Record A job
DART Tutorial

## Slide 6
<number>
DART Tutorial
Edit Parameter
Then this window will appear. Enter the value in this box you can reset to zero or preset to any value you wish
Then this window will appear asking for the password, Calibration should not be repeated unless the readings are different then the actual, therefore it is password protected to prevent unauthorised users from changing parameter calibration.
After clicking “Calibrate” as shown in the previous slide
After clicking “Reset” as shown in the previous slide
2
1
1
2
Reset and calibrate
Notes: <number>
Record A job
DART Tutorial

## Slide 7
<number>
DART Tutorial
Calibration (Analogue channels)
This slide Shows How to calibrate Analogue channels as Pressure sensors and Weight. For accurate Calibration you will need to input two reference points and the DART will calculate the Gain and Offset automatically
Before calibration make sure that the right parameter is selected
Current Raw value, the Reading from the H.W
EX:Take two different reference value Ex:0psi and 3000psi
The first Reference value is the actual physical value as the gauge EX: 0 PSI then enter zero manually in this box
When the first Reference value is applied click on this button in this case zero pressure applied
2
1
2nd Ref value: Is physical 3000 PSI enter the value manually here
Step
Step
When actual 3000PSI is currently applied press here
The GAIN and OFFSET will be calculated Then click “SAVE”
Notes: <number>
Record A job
DART Tutorial

## Slide 8
<number>
DART Tutorial
Calibration (Counter channels)
This slide shows how to calibrate counter channels as Depth, Pumping Totals, then further can be used for rates and speed
2
1
There is two methods for calibration, ether inputting K-factor if known, or using reference point as in analogue but one point only
Method
Method
If the K-factor is known as for ex: Depth then select “YES”
First Reset the parameter value to 0 as shown in slide 6
Then Input the Kfactor here Ex : 300 signal per one feet
In “Know K-factor” select “NO”
Taking 1 reference value will be enough. If we are calibrating any frequency signal as pumping totals.Then pump actual amount known using your unit, ex: 5 bbls, as a result number signals will be produce will be read on the raw value
Input the actual amount pumped the 5 bbls here
After pumping 5 bbl Click on “Do Ref”
Notes: <number>
Record A job
DART Tutorial

## Slide 9
<number>
DART Tutorial
Rates and speed calibration
After calibrating total counters (Pumping totals and depth) you can associate the to rates and speed
You can take the input the K-factor taking it from the total parameter relevant to it or you can do the method below to make the program associate it Rates with total counter and speed to depth
In “Know K-factor” select “NO”
After selecting “No” this list will appear asking user to select the relevant parameter to be associated with and taking its previous calibration performed
Click save when done
Then DART will automatically associate ex:(speed with depth and (N2 Rate with N2 total) and (fluid rate with fluid Total)
1
2
Notes: <number>
Record A job
DART Tutorial

## Slide 10
<number>
DART Tutorial
Parameters table
Parameters table is for setting up and editing parameters properties and graphs settings with every detail, it can’t be accessed during job recording, also if not ready to acquire
First go to “edit configuration” menu Then select “Parameters Table”
To identify channel type click on the cell then select from here
you can include or exclude parameters from graphs by entering in “Graph” column for each parameter its graph number that it will appear in.
For graph settings input the properties for each parameter in fields corresponding to it ex:parameters (name, unit, scale max,min, color,(no. format)) also this table shows the Gain and Offset it will be taken automatically from calibration so probably you will not need to change it from here
“BLANK” in Parameter “Name” means that it will disappear as if as it’s not there (not used
There is two other simpler way to edit parameter properties see next slide & slide 5
Notes: <number>
Record A job
DART Tutorial

## Slide 11
To change the parameter name at any time while recording, click the mouse left button while pointing the mouse cursor on the parameter name
Click here to exclude and move a parameter from this graph to another

## Slide 12
To change the parameter number format at any time while recording, click the mouse left button while pointing the mouse cursor on the parameter reading

## Slide 13
To change the parameter unit at any time while recording, click the mouse left button while pointing the mouse cursor on the parameter unit
To change the parameter graph Scale at any time, click the min. scale value to change it
And click the Maximum scale value to change it

## Slide 14
<number>
FACT Tutorial
PDF Report
Notes: <number>
Record A job
DART Tutorial

## Slide 15
<number>
DART Tutorial
Producing PDF JOB Report
Notes: <number>
Record A job
DART Tutorial

## Slide 16
<number>
DART Tutorial
Preparing Graph
Click here to include this graph to the report
Click close to proceed next
Notes: <number>
Record A job
DART Tutorial

## Slide 17
<number>
DART Tutorial
Weight VS Depth
And Wait for graphs to be loaded
Notes: <number>
Record A job
DART Tutorial

## Slide 18
FILL THE INFORMATION
THEN CLICK OK
ABC Petroleum
North Field
A-101

## Slide 19
Click here to plot the curves

## Slide 20
click here to add the graph to the report
If a specific time interval required, click here to zoom the time interval from the graph.
Move the mouse to the graph area, the cursor will change shape, click and hold the left button, then move the mouse vertically to the time required and release,for zoom
You can select time intervals manually by entering start and end time
This button to zoom out back to the full view time interval
After adding the graph to pdf you can repeat selecting time intervals for multiple graph pages
Click close to proceed next
N.B:The graph area will become blank, Use the PLOT CURVES button again to view curves

## Slide 21
ENTER THE NAME OF THE FILE

## Slide 22
THEN CLICK SAVE